## Slide 1
This presentation aims to summarize my experimentation over the last year or more,
with frequency generators particularly the world of GPS disciplined frequency references,
getting to solutions and measured results that satisfied me
and to explain the traps identified while assisting others
GPS Disciplined Frequency Reference
Traps for the Young and not so Young
The ’flavor of the month’ seems to be GPS locking of radio frequency.
To some, being “GPS locked” is enough, however the result is often not good
This subject is not new. I may be presenting solutions published before my examination
I have learned there are many errors and “wrong drum” in published material
George VK4AMG
VK4AMG@yahoo.com.au

## Slide 2
If you own one of these
You will know
It isn’t as easy as the magazines describe
GPS Disciplined Frequency Reference
Traps for Young and not-so Young Players
LOCKED !!
** Corrected * *

## Slide 3
Frequency Stability
  So my SSB / CW signal is clear and interference free
  So others optimum decode digital modes
  decode software is time and frequency critical
  So decoding other’s digital transmission is not degraded
GPS Disciplined Frequency Reference
Traps for Young and not-so Young Players
Frequency Accuracy
  So the other station can find me when I call on sked
and
  when operating in digital modes
      To be on specific frequency
      To be reported correctly in reverse beaconing
Within 100 Hz up to 1296MHz
Preferably better than 10Hz
For all operating conditions; Voltage, ambient temperature, power output, SWR, radio temperature, fan on/off

## Slide 4
Clock Gen
uP
‘Locked’
GPS Rx
Output
SPI Comms
Timer
Int
1PPS
Serial NMEA messages
OP1
OP3
Fo/8
32MHz
27MHz
Arduino Nano
uBlox Neo6
SI5351 module
GPS Disciplined Freq Gen
Proof of Concept Hardware
GPS DO Ref Gen Module

## Slide 5
Clock Generator e.g. Si5341
8 3.775MHz
Crystal Oscillator
EEPROM
SPI Bus Logic
PLL2
PLL3
PLL1
Phase Det Filter VCO
Switch Matrix
25MHz
725 MHz
29
Int
n/m
1 30.2MHz
24
6944
1048575

## Slide 6
Granularity due digital divider +- 1 Hz Si5351 far more for 4.4 & 13.4 GHz
1PPS interrupt latency 1uS produces 4 Hz
Correction latency e.g. Crystal drift of 10Hz in 10 secs 12 Hz
Frequency Count Resolution 1/8th freq 7 Hz
Errors in GPS Disciplined Freq Gen in 30.2 MHz Ref Output
NB* 1 Hz shift in Reference Frequency results in 50Hz at 1296 MHz operating frequency
40 ppb = 0.04ppm
Errors are additive e.g. >20 Hz at 30.2 MHz approx. 1 ppm 850 Hz at 1296 MHz
Include:

## Slide 7: GPS Disciplined Frequency Generator
Advantages
Simple low-cost hardware
Freq ranges 200kHz to 200MHz
138MHz to 4.4GHz
138MHz to 13.4MHz
Crystal frequency source for PLL
  Low cost
Disadvantages
Frequency accuracy limited by:
  Max Frequency of uP counter
  GPS and uP Interrupt Timing Jitter
  Frequency Count resolution
  Granularity of Freq Gen PLL and digital dividers
Frequency Stability limited by:
  Latency in frequency count
  Latency in applying corrections
  Crystal affected by ambient temp

## Slide 8
Performance GPS Disciplined Freq Gen 30.2 MHz
IC910H 27MHz OCXO Measure & Calculate 3.775MHz every sec Average error over 8 secs Applied 10 Secs
QRP Labs 27MHz Si5351 OCXO
Arduino Nano 3V 8MHz Neo6 1PPS

## Slide 9
Granulatity – Can’t be avoided – Use highest PLL frequency Carefully select integer and fractional divide
Interrupt latency - Processor selection – highest uP clock Code structure – avoid blocking functions Simple IRQ servicing code Frequency count over multiple 1PPS
Correction latency crystal selection – minimise frequency drift insulated environment heated environment OCXO
Frequency Counter - Select uP with async external input to timer - Avoid resetting timer and overflow - Count over multiple 1PPS periods e.g. 8 at 1/8th - Adopt Coarse / Fine Count & Correction Coarse 8 secs 1Hz Fine 80 secs 0.1Hz - Average Count or Rolling Average of Error
Overcoming Errors

## Slide 10
Correct by Fuzzy Logic not processor intense PID Controller
Consider count certainty in selecting dead band
Avoid compiled code structures known to be unfriendly to embedded uP e.g Variable arrays Select Case Indexed For Loops
Use specific frequency generator code not processor intense general case
Use device specific comms protocols updating only essential parameters
Avoid resetting frequency generator part unless necessary not necessary for small changes
Customise constants for specific hardware e.g don’t software correct for nominal crystal use close match considing operating conditions
Other Factors
It is not easy – Having GPS and a modern uP do mean that all will be well

## Slide 11
GPS Disciplined Freq Gen 30.2 MHz
IC910H 27MHz OCXO Measure 7.75MHz over 8 secs then Calculate Rolling error over 8 sec Applied 10 Secs
QRP Labs 27MHz Si5351 Synthesiser
Arduino Nano 5V 16 MHz Neo6 1PPS
This design minimised all errors, resulting in stability equal to the granularity of Si5351 (1Hz @ 30MHz)
More than a thousand hours of research, hardware and software development, testing, and modifications later
Is it good enough ? ? ?
Nearly Perfect but ….

## Slide 12
GPS Disciplined 10MHz Oscillator
IC7400 2m
MC145170 PLL
“Cal”
IC910H 23cm
Ref Gen
30.2 MHz
Freq Counter
IC7400 2m
Spectrum Analyser
32 MHz
Digital Modes PC
NMEA RMC Time Sync
Alternative Approach Away with Quantization Errors
Errors of the above GPS disciplined Frequency reference are avoided by Ref Gen or PLL locked to 10MHz Frequency Standard
Frequency and time
A whole shack solution

## Slide 13
10Mhz Locked Frequency Generator
uP functions minimised
  Initialise Clock Gen
  Detect 10MHz Standard
  Enable output of Clock Gen
  Switch off supply to radio ref
Similar hardware, no GPS, depending on external 10MHz Standard to avoid quantisation errors
May be applied to most modern radio designs using an Internal Frequency Reference (incl SDR)

## Slide 14
Rubidium / Caesium Frequency Standard 0.001ppb 0.001ppb Used $500++
GPS Disciplined Oscillator 0.01 ppb 0.01 ppb e.g. Trimble Thunderbolt Used $500 G7TBL GPSDO New $300
Oven Controlled Crystal Oscillator Not Calibrated 0.5 ppb e.g. Morion MV89 Used $50 IsoTemp-131 Used $50
Temperature Compensated Crystal Oscillator Not Calibrated e.g. SiTime 5356AI 0.1ppm New $75 IQT LF TVX009900 0.2ppm New $20 CTS 580L100X2 0.25ppm New $20
Voltage Controlled Crystal Oscillator Not Calibrated e.g. AVVV -10MHz I25 25ppm New $20
Options for 10MHz Reference
Accuracy Stability / Day Price
The lower cost options may be Manually Calibrated or GPS Disciplined
And must be insulated from environment or otherwise temperature controlled Temp compensated not enough.

## Slide 15
uP
‘Locked’
GPS Rx
Output
Timer
Int
1PPS
Serial NMEA messages
5MHz
10MHz
TCXO 10MHz IsoTemp-131
2
PWM
Serial
GPS Disciplined 10Mhz Oven Controlled Oscillator
Le
Level Shifter
0 -> 8V
Arduino Nano 5V 16 MHz
Neo6
Configure for NMEA RMC only

## Slide 16
GPS Disciplined 10MHz IsoTemp-131 VCXO
OCXO excellent accuracy +- 20 ppb Uncorrected
Fine Pulling Range +- 4ppm
Supply 12V at less than 12W < 1.5W quiescent
Control range 0 – 8 V dc PWM Level Shifter or AD7243
Sine wave output at +8 dBm
Disciplined Accuracy < + - 2 ppb
Fast turn-on Use last correction (EEPROM) after warm-up
Little or no discipline necessary Use not disciplined - occasional calibration

## Slide 17
GPS Disciplined 10MHz Morion MX89 TCXO
Output 10MHz sine +7dBm
Stability < 0.0002 ppb
Power Supply 12V @1.5A warmup 350mA Quiescent
Warmup 15 minutes
Pulling Range 25 ppm
Pulling Volts 0 -> 5V
+- 1 Count on HP5384

## Slide 18
Clock Gen
Microprocessor
‘Locked’
GPS Rx
Output
SPI
Timer
Int
1PPS
OP1
OP3
Fo/8
30.2MHz
27MHz +-
GPS Disciplined VCXO in Freq Gen Also away with Quantisation Error
VCXO
PWM
Granularity issues avoided
Frequency count, VCXO control, & stability controllable

## Slide 19
Improved ‘Prog Rock’ - GPS Disciplined ASVV VCXO
Arduino Uno QRP labs “Soft Rock” Si5351 module on QRP Shield
27 MHz VCXO mounted on Arduino Uno shield is controlled by PWM from Timer 1
PWM is filtered via two stage RC filter 27ms and 270ms time constant
One tenth of DC control signal is mixed with split 3V3 reference signal
Network provides +- 1.5 kHz control range with 216 -1 steps (1 ppb)
Si5351 OP3 provided <5MHz to Timer 1 for Counting and PWM generation
16 MHz Arduino Uno AT328n Timer measures 5MHz with excellent repeatability
Eight successive frequency counts are added to calculate frequency error
Rolling average of error is calculated over eight further seconds
This error is applied by fuzzy logic to PWM duty cycle
Thirty second settling period allowed between corrections
Less than five minutes is required to settle output within 10Hz
Another 5 to 10 minutes settles the output frequency to better than 1 Hz accuracy
NB: ASVV VCXO is affected by ambient temperature changes Typically + - 25 ppm
Temperature, counting, correction errors need to be considered and controlled

## Slide 20
OCXO Si5351 Synthesizer
This module aims to improve on the ‘Prog Rock’
Provides temperature-controlled oven
Isolate load from crystal oscillator
Improve power supply regulation and isolation
Claims better than 1 ppm stability (enough ???) over 2 minutes WSPR transmission
In QRP Lab application of WSPR beacon frequency is not measured during transmission
Frequency discipline may use GPS 1PPS over long term
I have not attempted to measure of compare the QRP Labs modules for frequency accuracy Comments on potential error for other GPS Disciplined Si5351 and AT328P modules apply
Combining “Improved “Prog Rock” with QRP Labs OCXO should provide a 1ppb accuracy & stability

## Slide 21
Conclusions
Modern Digital modes require high frequency accuracy and stability as well as accurate timing
Digital Frequency Generators have quantisation limitations which cannot be overcome
Combination of 10MHz “Standard” feeding PLL or Frequency Generator provides good result
GPS disciplined VCXO may be combined with digital Frequency Generator to provide a good result
10MHz “Standard” or VCXO with Freq Gen must be insulated from temperature changes
For more info on Silicon Chip Freq Gen – send email with info / photos of your attempt VK4AMG@yahoo.com.au